# University of Virginia Health System

# Medical Laboratories

# “Quality You Expect, Service You Deserve”

# LABORATORY MEDICINE UPDATE

March 5, 2014

Change of PTH Assay

On February 13, Abbott Diagnostics notified UVA of a recall of their Architect reagents for measurement of parathyroid hormone (PTH). The recall cited a 13%-45% increase in results with reagent lots used at UVA since January 2013.

UVA Medical Laboratories has taken the following steps: (1) On February 13, the laboratory halted use of the Abbott PTH reagent,  (2) implemented an  alternative method (Siemens Immulite assay) on February 18-20, and (3) completed analysis of all (~200) samples in the laboratory on February 21 with the Immulite assay.

The assay in use now (Siemens Immulite) had been used at UVA until May of 2011 when it was replaced with the Abbott assay. At that time, the results of the two methods agreed closely, with a mean difference of 5 pg/mL. Comparison of the results of the two assays now shows a mean difference of 30%-32%, consistent with the manufacturer’s recall indicating an increase of 13%-45% in the Abbott results.

In view of the potentially misleading results of the PTH assay during the last year, repeat PTH testing will be performed as needed at no charge. To arrange no-cost retesting, contact Dr Haverstick at dmh2t@virginia.

Occult Blood Specimens

When providing patients with the Occult Blood Kits for specimen collection in the home, please attach a clinic sticker from the patient’s visit to the box.  This will provide the lab with the correct ordering location and physician and prevent any delay in testing while the request is investigated.  This will also assure that the results are sent to the correct physician.

Specimens to Be Sent to the State Lab or the CDC for Testing

At the request of Virginia Department of Health, Clinical Microbiology, and Hospital Epidemiology, the process for sending samples for testing to the Virginia State Lab (DCLS) or any CDC Laboratory has been formalized.  The Clinical Microbiology Director on-call (pic 1221) will coordinate with Virginia Department of Health to be sure that such testing is considered warranted by them.  If a specimen is submitted to the Labs without prior authorization from the Microbiology Director on-call, the labs will notify pic 1221 who, in turn, will contact the requesting physician for clinical symptoms and to ensure that the appropriate paperwork is completed.

The tests below list are representative of such requests; the list is not all-inclusive.

Amoeba			Anthrax		Botulism			Hantavirus			Diphtheria		Measles - IgM and/or PCR

Mumps - IgM and/or PCR	Tularemia		Viral Hemorrhagic Fever

Brucella			Smallpox		Vaccinia

Yellow Fever			Chikungunya virus

Testing for Arbovirus does not require authorization.  However, a specific form, available on the UVA Medical Labs web-site, is required:

http://www.healthsystem.virginia.edu/pub/medlabs/lab-handbook-test-directory/labgeneral/requisitions/state-lab-submission-forms