# University of Virginia Health System

# Medical Laboratories

# “Quality You Expect, Service You Deserve”

# LABORATORY MEDICINE UPDATE

April 27, 2015

Discontinuation of In-House Platelet Factor 4 Antibody Test

This past summer, a platelet factor 4 antibody test was implemented for use in the evaluation of patients with suspected heparin-induced thrombocytopenia (HIT).  The test’s performance has proven to be unacceptable. When all 151 results of the test were compared with the results of the reference standard serotonin release activity (SRA) test, 57 were false positives and 6 of the 7 samples that were SRA-positive were negative.  As a result, the test will be discontinued, effective April 29, 2015.

The Anticoagulation Committee is updating the guidelines for HIT, which will include additions to the in-house testing menu.  In the meantime, the following is advised for patients in whom HIT is suspected:

- Calculate a 4T score http://www.healthsystem.virginia.edu/docs/manuals/guidelines/cpgguidelines/02-clinical-practice-guidelines/hit-hitts-adult/?searchterm=HIT

- If 4T score > 4 , send-out SRA

- Consider Hematology consult, if needed

NP Swab Samples Accepted for B. pertussis PCR Testing

As of April 23, 2015, the Clinical Microbiology and Molecular Diagnostics Laboratory will now accept nasopharyngeal (NP) swabs for B. pertussis testing. Acceptable swab types are as follows:

- mini-tip polyester or rayon aluminum swabs

- flocked mini-tip swabs

Required specimen: NP swab must be placed into a sterile tube containing 3.0 mL M4RT or UTM-RT at the time of collection. A single sample submitted will usually be sufficient for B. pertussis PCR and Viral Respiratory testing.

Testing Schedule:  B. pertussis testing will continue to be performed once daily.

Please contact the Clinical Microbiology Director on call at PIC 1221 with questions.

Discontinuation of In-House Phosphatidyl Glycerol Testing

Due to discontinuation by the vendor of reagents for phosphatidyl glycerol testing, Medical Laboratories will cease testing on May 1, 2015.  Testing of amniotic fluid for Lamellar Body Count will remain in place.

Thyroid Testing and the Medicare Patient

Thyroid function studies (defined by Medicare as: T4, Free T4, TSH, and T3 Uptake) are used in delineating the presence or absence of hormonal abnormalities in the thyroid and pituitary glands. They are used to define hyperfunction, euthyroidism, or hypofunction of the thyroid.  Routine screening (V70.0) is not a covered service for Medicare patients, but testing may be considered reasonable and necessary to:

- Distinguish between primary and secondary hypothyroidism

- Confirm or rule out primary hypothyroidism

- Monitor thyroid hormone levels in patients with goiter, thyroid nodules or thyroid cancer

- Confirm or rule out primary hyperthyroidism

- Monitor therapy in patients with hyper or hypothyroidism

Thyroid testing may also be medically necessary in patients with metabolic disorders, malnutrition, hyperlipidemia, certain types of anemia, psychosis, menstrual disorders, skin conditions, etc. along with a wide variety of signs and symptoms, including alterations in consciousness, malaise, etc.

A few examples of the more common diagnosis codes are 242.00, 244.9, 250.00, 401.9, or 780.79. If these or other codes are not applicable, please discuss with your patient and have them sign the Advanced Beneficiary Notice (ABN) for billing purposes.

Testing may be covered up to two times a year in clinically stable patients; more frequent testing may be reasonable and necessary for patients whose thyroid therapy has been changed.

As a reminder, the preferred algorithm for assessment of thyroid function is to first perform a TSH (Thyroid Stimulating Hormone) analysis.  If the value falls outside of the range 0.45 to 4.5 mIU/L, then a free-T4 (free thyroxine) test should be performed.  The University of Virginia Medical Laboratories offers this reflex algorithm as TSH Reflexive in all ordering systems.  Please contact the Clinical Chemistry attending (pic 1657) or resident (pic 1267) on-call if you have any questions.

Samples for Ammonia Analysis “on-ice”

Ammonia is a natural metabolite that continues to be formed in whole blood after a sample is drawn.  This process is slowed when the blood sample is chilled after drawing.  Thus current guidelines indicate that samples for ammonia analysis should be submitted on ice.  Effective January 15, 2015, samples will be rejected if this requirement is not met.

A reminder:  On-ice means placing a bag of ice into a second bag that contains the sample.  These should be placed in a third bag prior to submission to the Core Laboratory.  Do not place the sample in the bag with ice as this causes the label to disintegrate.